Responsibility

Line Leaders - carry out routine tests, and fail-safe checks at the defined frequency. FST - complete Daily verification checks of the detection system. Audit the Line Leaders against the testing procedure. Investigate rejected products. Verification of the detection equipment after repair work, machinery movement or new equipment installed. Ensure routine testing and failsafe mechanism are working after engineering work on the x-ray. A handover must be carried with FST and recorded on the line check sheet. Area Manager - Ensure Line Leaders are trained and competent. Check & countersign the completed records during the shift. Engineering - Resolve / repair detector failures and malfunctions. Ensure routine test and failsafe mechanism working. Engineers must complete the handover document in the check sheet before leaving the line. Complete PPM checks as per the defined frequency. Technical - Ensuring the X-rays are calibrated as per the calibration schedule. Compliance/ HACCP team - To ensure quarterly CCP audits are being carried out. calibrations/ validations are completed as per the schedule and following the movement of equipment or the introduction of new products.

Frequency

Start up: This must be completed at start of the shift and completed again at the end of the shift. Also this check must be completed after engineering work on the line, fire alarm or major breakdown Routine Start of each product run, hourly thereafter and at the end of each product run. After any breakdowns or maintenance. Following breaks. Before restart after an emergency evacuation.

Procedure

Scope To reduce the risk of contaminated product reaching the end consumer and achieving the best detection of foreign bodies. Overview of Procedure: The effective X-ray of product is defined as a critical control point in the process and is essential in demonstrating a diligent approach to foreign body prevention to our customers, and enforcement authorities. It has been identified within our HACCP study as being a key critical control point and an area, which is most likely to identify product contamination if this procedure is followed effectively. If X-Ray detection equipment is moved within the factory, then the equipment’s performance and operation must be re-validated before being used. When new products are introduced, the detection equipment’s performance and operation must be validated. This is an integral part of the HACCP process and any X-Ray detection equipment must not be used until this process has been completed and recorded by the HACCP team. Principles of X-ray inspection How X-rays Inspection Systems Work An x-ray inspection system has three main components: a device to generate x-rays, a sensor to detect contaminants and a high-speed computer to process sensor data The Process X-rays are generated by a special glass tube that is cooled. When high voltage is applied to that tube, electrons are generated. The electrons then strike a target which emits a stream of x-rays that are focused and reduced to a fan beam. The fan beam passes through the product before entering a sensor. The sensor then converts the x-ray signal into a form where it can be passed to a computer for processing and detection. Detection X-ray systems detect contaminants by density difference. As an x-ray penetrates a product, it loses some of its energy. A dense area such as a contaminant will reduce the energy even further. As the x-ray exits the product, it reaches the sensor - an array of photo-diodes that precisely measure the x-ray energy that reaches them. The sensor converts the energy signal into a grey-scale image. A high-speed computer scans the sensor at thousands of times per second and stores the individual signals. It compares them to a dynamic value that delivers the maximum level of sensitivity and minimum level of false readings. A contaminant is highlighted as a darker shade of grey than the product. Effective x-ray inspection depend on the sophistication of data analysis Critical Limits As the detection of contamination through X-Ray in finished product is classified as a CCP (Critical Control Point), critical limits have been set. Critical Limits are a set of criterion which separates acceptable limits from unacceptable. All lines fitted with X-ray in High Risk are set up to the critical limits shown below S/S Ball; S/S Wire; Aluminium; Glass. 1.0mm; 0.5 x 5.0mm; 4.0mm; 4.0mm. Critical limits for ProSeal Line X-ray and specific M&S frozen retail products through frozen X-ray (Twin pack turnovers, pastry blocks & twin pack pies) S/S Ball; Aluminium; Glass. 1.0mm; 4.0mm; 4.0mm. Duck Cassoulet Gastro Meal on ProSeal Line S/S Ball; Aluminium; Glass. 1.0mm; 5.0mm; 4.0mm. Bulk Frozen Products S/S Ball; Glass; Ferrous. 1.5mm; 8.0mm; 6.0mm. ## X-ray operation Before the X-Ray can be used, startup checks are required to be completed to ensure that the machine will start correctly. - Open Both Panels on the front to ensure that there is no debris, ensure that there has been no damage occurred to any of the skirts. - Before starting the machine, ensure that the skirts located on the in feed are tucked into the X-Ray machine to ensure that they do not block off the sensor and prevent the machine from working. - On start-up of the X-Ray the black side switch must be turned on to allow the X-Ray machine to warm up and the screen to load. Once this has been completed and the screen is ready to use, the small key can be inserted into the front, without this key in place the machine will not run. - Turn on the X-Ray using the screen and pressing the Green X-Ray Button, this will start up the X-Ray warm up sequence. Once this button has turned red, and the three lights on the Front, left side and right side have turned red this X-Ray is on. - The programme for the product which is going to run needs to be selected, this is to match the programme selection on the Checkweigher machine. - The Green Button which contains the word Conveyor can now be selected and will turn Red to highlight that the conveyor is now running. - When operating the X-Ray Machine, the start-up checks on the sensors are required to be completed as detailed below - If the X-Ray does stop for any reason then the X-Ray lights on the Front and sides will go off and the button on the front will turn Green, this can be pressed to re-start the machine, but it will need to warm up, its progress will be seen on the screen and the button will turn Red and the lights on the Front and Sides will come back on when the machine is ready to run, and the conveyor can then be turned on. Testing Procedure Checks are completed on the X-Ray as a minimum at start of production run, hourly and end of Production run. At varying levels of access, the Factory Technical Development Manager, Food Safety Technologist, FST and Engineers have access to the menu that allows adjustment to sensitivity of the X-Ray. THE X-RAY MUST BE CHECKED AT THE END OF THE PRODUCTION RUN TO VERIFY THAT ALL PRODUCTS SINCE LAST TEST ARE SUFFICIENTLY DETECTED THIS INCLUDES END OF SHIFT. # Routine Testing: # IMPORTANT: TEST MODE MUST NOT BE USED WHEN CARRYING OUT THE ROUTINE CHECKS To ensure that the X-Ray system is fully operational and effective a series of tests must be completed at the defined frequency. The rejects from the box and captured images are saved in a memory card in the x-ray machine which downloaded for examination to ensure that the test pieces are being picked up and the size level it has achieved which is within the critical limit. If any test pieces are not detected at the correct level of sensitivity, then all products since the last good check are to be placed on hold, FST must be informed. Following conformation from FST that the machine is working correctly. The reject bin should be checked that it is empty before any test is completed on the X-Ray. NOTE: Any product found in the reject bin must be investigated as per the reject procedure below. This can be aided by using the DATA button on the X-Ray, and selecting the button which says NG image, this will bring you up images which have been identified as Not Good (contain a potential foreign body) to identify the area of the potential contamination. # If the test piece is not detected by the x-ray, the test piece will not reject into the reject bin and must be recovered immediately to avoid the test piece being packed and distributed. # Frequency The X-Ray function must be routinely tested at the following intervals; - Start and end of every product run - After breaks - Following maintenance - On restart following an emergency evacuation - Hourly checks Note: The frequency of testing during production runs must not be greater than hourly intervals. This frequency of testing allows for the identification, recall and testing of all products which have passed through a failed system since the last successful test. Test Packs Test packs must be made up from products that have successfully passed through the X-Ray system and must be clearly marked and labelled with the METAL TEST label to prevent the samples being inadvertently distributed. The test packs must match the finished product e.g Pack 12 sausage rolls rather than individual The metal test label must include the following information: Line Leader name, date and time of test pack generation, test piece type and test piece size. High Risk The samples are taken either from the sample rack in the product chiller or if at the start of a new week, then four new samples will be taken from the line after successfully passing through the X-Ray detector. Low Risk The test samples are taken from the storage. Due to the constant rejection, samples can become damaged. So, will need to be periodically changed for a new sample. Samples also must be checked that are still in date and any out of date products disposed of. The new samples will be taken from the line after successfully passing through the X-Ray detector. Test pieces Each line has a locked box that contains the X-Ray test pieces. The box must be locked at all times. In High Risk, the locked boxes are checked by the Line leader/Admin upon issue and return to ensure that all the test pieces are present and not damaged, the issue and return of the test pieces are logged on form ref: FSR048b In Low Risk, the locked boxes are stored in the Frozen Pack & Proseal line cages. Test pieces and packs must not be stored on top of the X-ray equipment or in the locked reject bins; this can affect the performance of the detectors. Recording of test results When recording the test results, the line leader must record these on the relevant CCP record sheet, the certificate No of each test piece used should also be recorded. - Low Risk X-ray Record Sheet: Ref FSR036c - High Risk X-ray Record Sheet: Ref FSR048a - Proseal X-ray Record sheet: Ref FSR036f Routine Testing Procedure Start & End of every product run - minimum Hourly. Testing must be carried out by placing all four test packs/units onto the line, then passing them through the X-Ray consecutively at normal product placing. Fail Safe Start Up, Start & End of every product run See diagram for each position. The four test pieces should pass through the aperture as close to the centre as possible. Direction of finished pack/unit along belt Top of pack Test piece 1 Top of pack Test piece 2 Top of pack Top of Pack Test Piece 4 Test Piece 3 - Take four full packs containing product or four units which have already passed successfully through the X-Ray, (either from the sample rack or if start of shift pattern take four new samples) - Securely tape each test piece to the top of a full pack/unit in positions 1, 2, 3 and 4 (as above) clearly mark as metal test packs using the Metal Test labels provided. Aluminum/Ferrous and Glass test pieces must be placed in either leading or trailing position. This is because for the products with uneven fill the combined density of the contamination and product depth is lowest in this area making detection more difficult. Stainless Steel Ball and Stainless-Steel Wire must be positioned over the centre of the product. - Pass the packs/units through the X-Ray consecutively (one straight after the other as per the normal spacing between packs/units when running during production). Allow the pack/unit to travel down the belt and be deflected into the collection bin – DO NOT REMOVE PACKS/UNITS BEFORE THEY GO IN THE BIN. The pack/unit should be rejected & fall into the reject bin (if not refer to Actions below). - Once test is completed, the test pieces must be removed and placed back into locked box. Samples must be returned to storage or the Freezer to re-use daily. - Record the results on the relevant area record sheet as P for Pass and F for fail. Memory Test: Start & End of Shift - A memory test needs performing at the start and end of the shift. - Ensure that the packs/units to be used for the test have successfully passed through the detector head before being removed from line to set up as test packs/units. - Securely tape each test piece to the top of a full pack/unit in positions 1, 2, 3 and 4 (as above) clearly mark as metal test packs using the Metal Test labels provided. Aluminum/Ferrous and Glass test pieces must be placed in either leading or trailing position. This is because for the products with uneven fill the combined density of the contamination and product depth is lowest in this area making detection more difficult. Stainless Steel Ball and Stainless-Steel Wire must be positioned over the centre of the product. - Allow the packs/units to travel down the belt in the order shown on the diagram below. X-RAY Product test leading edge Std product Product test Middle pack Product test Std product Product test Std product ALL the test packs/units must be deflected into the reject bin! DO NOT REMOVE PACKS/UNITS BEFORE THEY GO INTO THE BIN. - Once test is completed, the test pieces must be removed and placed back into locked box. Samples must be stored. - Record a P for Pass and an F for fail on the relevant record sheet - If the test fails alert the FST/ Technical team who will advise on a suitable course of corrective action to be taken. Fail-safe Checks Failsafe checks will be conducted at the start of each shift, End of shift, major breakdown or engineering repair - Reject Bin Full - This is an automatic belt stop safe system that activates when the reject collection box becomes full. This is to stop the bin being so full that packs/units can’t be rejected into the bin. In the event of the reject bin becoming full of product a sensor will detect this and stop the belt. To test if this mechanism works, turn the sensor cover until it covers the bin full sensor. The “Bin Full” light will illuminate, the alarm will sound, and the belt will stop on the next reject pack/unit (pass through a pack/unit with a test piece attached as above). Record the result on the relevant record sheet. - Reject Confirmation – This is an automatic belt stop safe system that captures contaminated products that have not successfully entered the reject bin. All contaminated products should be rejected into the bin. When packs/units go in the bin they pass across a sensor. If the sensor is not passed the line will stop. For a successful rejection to occur the system has a sensor that detects whether a pack/unit has been effectively diverted to the secure bin. Testing MUST be carried out by passing a test pack/unit down the line while temporarily overriding the rejection system (by either using a key switch held by an authorised person or via selecting the reject inhibition option on the control screen) and observing that the reject mechanism does not operate, and that the conveyor belt then stops. The “verification failure” light will illuminate or Reject confirmation error message will be displayed the alarm will sound and the belt should stop with the product not leaving the end of the conveyor. Frozen packing belt stop timing product will stop on the grey transfer rollers outside of the machine before passing through the wall hatch. Record the result on the relevant record sheet - Reject Bin Door Status Monitor – To prevent unauthorized access to rejected products and reduce the risk of the product being sent out to the customer the bin door must be locked. To prevent the systems being able to run with the door open an electronic interlock device is fitted to the reject bin door. To test this, start the program and unlock and open the bin door. The belt should stop, and the alarm will sound within 30 seconds (ideally around 10 seconds). Test this and record the result on the relevant record sheet - Air pressure failure - This is a means of checking that the air failure switch is working. This is an automatic belt stop safe system that captures contaminated products that have not successfully entered the reject bin due to air pressure failures. In the event that the air pressure is lost, the reject mechanism will be unable to work. There are two types of the reject set up used on site: 1. Reject mechanism with the reject bin attached to the X-ray – the air pressure is connected directly to the X-ray machine (Frozen Packing Line, Encroute line and High Risk Line 3). This will immediately reject product at the X-ray machine. 2. Reject mechanism with the reject bin attached to the conveyor belt after the X-ray machine (High Risk Lines 1, 2,4 and 5) where X-ray uses the Dacs (Checkweigher) air pressure to reject the product. Both reject mechanism are tested by starting the program and isolating the air supply – the belt should stop and the alarm will sound. Test this and record the result on the relevant record sheet. - Audible alarm and light – To ensure that when the X-Ray detects a contamination both the alarm and light should sound/ Light up in all the above events, record the result on the relevant record sheet THE X-RAY MUST BE CHECKED AT THE END OF THE PRODUCTION RUN TO VERIFY THAT ALL PRODUCTS SINCE THE LAST GOOD TEST ARE SUFFICIENTLY DETECTED. THIS INCLUDES END OF SHIFT. Every test pack/unit must be detected and rejected for the routine/ fail safe tests to have been successful. If any part of the routine/ fail safe test fails, production must be stopped immediately. Inform the area FST and engineering department contacted to allow necessary investigations and corrective actions to be completed. It is not acceptable to repeat the tests until the packs/units are rejected successfully. Do not attempt to alter the settings or sensitivity unless authorised to do so. Corrective Action If a routine, fail safe or memory test fails against the critical limits, contact the area FST who will quarantine the entire product on that line since the last satisfactorily completed test. If required, the engineering team will be contacted to investigate and rectify the issue. Upon completion, all the start-up and fail-safe checks must be completed before production can restart. These re-tests must be fully documented on the relevant record sheet. Search head failure If the X-Ray goes into fault mode all products that remain on the belt must be removed and passed back through the X-Ray. Following the activation of a fault with any of the X-Ray/ rejection system during normal production which results in a belt stop all the start-up validation and fail-safe tests must be successfully completed before production restarts. Note: Ensure that any packs with test pieces attached are retrieved which may have entered the collation area (all packaging must be sanitized upon re-entry into high risk). Ensure all test packs are returned to the storage rack in the chiller area once the tests have been completed. Quarantine of product All products (both fresh and frozen) that has passed through the process since the last good check must be identified, isolated, placed on quarantine and re-x-rayed through a working X-ray. Rejected Product Any product found in the reject bin must have a METAL REJECT sticker placed on immediately to make sure that it can be clearly identified as a potentially contaminated pack/unit. If rejection of product occurs record number of rejected packs/units on the relevant record sheet. The number of reject packs in the bin must be checked against the number of rejected packs by the x-ray machine. All rejected products must be tested back through the x-ray 3 times (in different orientations if possible). M&S Product: must be disposed - including Bulk products. Other customers product: if all 3 re-tests come back clear the product can go to customer. If any of the 3 retests are rejected the product must be passed to FST for inspection and wasted. All Reject bins can be only accessed by Authorised personnel. All reject bins must be locked, keys must be kept in the locked box and can’t be left in the lock. Reject confirmation errors All failed rejection belt stop alarms that are triggered outside of the failsafe tests must be treated as a rejected product. The reject confirmation failure will trigger a belt stop, alarm and error message or “verification failure” light display. The product/pack that has been identified as containing a possible foreign body/contaminant will be stopped within the downstream flow of the X-ray detection area. If the line is fitted with an in-line guarded checkweigher (High Risk 1,2,3,4 and EnCroute) then the rejected product will be the product immediately after the reject arm. This product must have a ‘Metal Reject’ sticker applied and follow the above process to investigate if the product requires investigation by FST. Frozen packing X-Ray does not have any guarding or equipment downstream of the X-ray reject bin. In the event of a Reject confirmation error and belt stop on Frozen packing then the last 4 printed retail packed products or last 1 bulk packed product must be removed from the grey transfer rollers just after the X-Ray outfeed. These products must be treated as containing a possible foreign body/contaminant and must have a ‘Metal Reject’ sticker applied and follow the above process to investigate if the product requires investigation by FST. If you are unsure of the specific product that failed to reject, then all products in or downstream of the X-Ray must be treated as a ‘Metal Reject’ and cleared from the line. Ensure that the reject belt and detection area of the X-Ray is clear of any products and that the reject arm does not have any obstructions or product caught behind. Corrective action after reject confirmation error must be recorded on production paperwork with the number of products/packs marked as ‘Metal Reject’ related to the reject confirmation event. If more than one reject confirmation error occurs during a single product run or within 1 hour engineers, Line Manager & FST must be informed and Engineers must investigate the issue. Multiple Rejects (2 or more) Multiple rejects of a product within a short period of time must be investigated immediately and all information including corrective actions must be recorded on the investigation. refer to the Foreign body procedure for actions M&S CODE OF PRACTICE Maintenance A full test is completed following any repairs, maintenance or adjustment of the detection system before the line is returned into production. Training All line leaders and engineers who are carrying out X-Ray checks and repairs must be trained against this procedure to ensure that they fully understand the principles, effective operation and corrective actions required to prevent product contamination occurring in our products to protect the customer. All new Line leaders must take the competency test to prove their understanding of this procedure (FSR024a) before being allowed to run the production line.

References

- Highlighted in bold above

| History of Change |  |  |  |
| --- | --- | --- | --- |
| Date | Previous Version No. | Reason for change | Authorised |
| 25.02.15 | 01 | Changed Validation to Verification following BRC audit observation | S Hawkshaw |
| 26.03.15 | N/A | Changed wording slightly on pages 3, 5 & 8 to avoid confusion of the requirement. | S. Hawkshaw |
| 11.11.15 | 2.1 | Removed the need to complete higher level compliance checks – photogating sensor and Drive roller as these are not part of the customer COP for X-ray | S. Hawkshaw |
| 11.10.16 | 3 | Updated following internal audit non- conformance | K Kolomyjec |
| 03/08/17 | 4 | Document amended, test pack label must include the following information: Line Leader name, date and time of test pack generation, test piece type and test piece size. Addition to the box on each line for x-ray test pieces to be locked at all times. Test packs to be changed weekly in High Risk. | S. Bolton |
| 19/12/2017 | 5 | Document amended, reject bins section renamed to Rejected Product and “Any two or more rejects found in the reject bin must be retested immediately:” added to it. Reject Pusher arm failure removed from Fails safe checks, Reject verification changed to Reject Confirmation, | S Bolton |
| 09/01/2018 | 5.1 | Document amended, Test pieces placement updated for both memory and consecutive test following the guidance received from the x-ray system provider. | S Bolton |
| 19/06/18 | 6 | Critical limits updated for Dinky sausage rolls. X-raying of individual units in Low risk frozen packing added. Details regarding reject mechanism for air pressure failure added. | C. Manoj |
| 05/04/19 | 6.1 | Added 15 pack Cocktails to the increased test stick size, Updated the format and removed the “b” from the title, changed Metal detection to XRAY detection | A.Singh |
| 13/05/19 | 6.2 | Added reference to the Encroute Line | A.Singh |
| 10/06/19 | 6.3 | Changed the word Cause to Identify on page 1 & Removed freezer reference “(LR)” From Page 5 | A.Singh |
| 18/07/19 | 7 | Added Bulk Frozen Sticks, added Metal Pens to be tested on Product and Added Rejects to be Retested in Different orientations. | A Singh |
| 05/09/2019 | 8.1 | Added Bulk Frozen Specifications for X-ray Stick Sizes | A Singh |
| 18/09/2020 | 8.2 | Typo on the Glass Spec stated 4.00mm should be 8.00 on Frozen Packing Line – Documentation Is Correct | A.Singh |
| 14/05/21 | 8.3 | Bulk x-ray reject actions added | C. Jones |
| 03/03/2022 | 9 | X-ray test mode removed when completing x-ray checks. X-ray checks must not take place in test mode. | C. Jones |
| 29/06/2022 | 9.1 | Update test piece critical limit tables to reflect current and seasonal products | L. Moan |
| 04/11/2022 | 10 | Updated Responsibility and Frequency part to include handback procedure, fails safe activity and updated products of note with specific test piece sizes & addition of guidance for Reject confirmation error and belt stop line clearance. Inserted extract of M&S COP relating to two or more metal finds in same production run or 24 hour period. | R.Tettey |
| 14/02/2023 | 10.1 | Update reference to FSR036b Pens and Plasters X-Ray Check on Quor. | G. Birkens |
| 22/12/2023 | 10.2 | If the root cause identified the foreign body as coming from a supplier within a raw material, this must be collected and returned to the supplier to allow supplier investigations to take place. | G. Birkens |
| 08/01/2024 | 10.3 | Removal of reference to near miss and non-conformance log. | G.Birkens |
| 30/04/2024 | 10.4 | Addition of checking reject bin count against rejected product count on X-ray. Removal of plasters under x-ray detectable consumables. | G. Birkens |
| 13/04/2025 | 10.5 | QA changed to FST, and Senior QA changed to Food Safety Technologist | R. McDonnell |
| 30/05/2025 | 10.6 | X-ray check of factory pens removed & added to FB procedure. | S Ford |
| 21/08/25 | 10.7 | Removal of test pieces sizes for delisted products beef and bone marrow and 15pack cocktail | R.Smith |
| 10/10/2025 | 11 | Added test piece sizes for Duck Leg on Poseal Line Addition of proseal paperwork Clarified rejected product section, removed detail on FB corrective actions as detail is in FB procedure. | R.Smith |
| 17/11/2025 | 12 | Removed the requirement for S/S wire test pieces for Low Risk. Amended where low risk test piece boxes are stored to align with current practice. Added requirement to record Cert No of test pieces used on the record sheets. | S.Ford |